## Reference
Reference Spec
Manufacturer | LUTRONIC
Model | enCurve
Operating frequency | 27.12MHz
HF power | 10 ~ 300W
Treatment Mode | continuous
Power consumption | 1300 VA
Power connection | AC 220 V
Mains frequency | 60 Hz
Dimensions (W x D x H) | 475 x 758 x 1133 mm
Weight | 67 kg

## System Budget
DDS |  | AD8051 |  |  |  |  | LPF |  |  | Measure |  | LT1210
Vpp | Vrms | R34 | R33 | Gain | Vpp | Vrms | IL | Vpp | Vrms | Vpp | Vrms | R4 | R7 | Gain | Vpp | Vrms | T1_OUT- | T1_OUT+
0.6 | 0.2121 | 4990 | 1000 | 5.99 | 3.594 | 1.2705 | 6 | 1.797 | 0.63524 |  |  | 680 | 680 | 2 | 7.188 | 2.541 | 3.594 | 3.594
0.6 | 0.2121 | 470 | 47 | 11 | 6.6 | 2.3331 | 6 | 3.3 | 1.1665 | 2.8 | 0.9898 | 680 | 100 | 7.8 | 21.84 | 7.7204 | 10.92 | 10.92
[7 empty] | 9750 | 9850
[7 empty] | 162.5 | 164.1667 | 326.6667
[7 empty] | 2.7083 | 2.7361 | 5.4444

## Power
Power |  |  |  |  | LED
SMPS | 300 | 600 | 600 | W | V | 12 | 12
Voltage | 48 | 48 | 48 | V | V_LED | 1.9 | 1.9
Current | 6.25 | 12.5 | 12.5 | A | R | 2 | 10
Pin Current | 3 | 3 | 7 | A | I | 5.05 | 1.01
Pin # | 2.0833 | 4.1667 | 1.7857 | ea
[10 empty] | MCLK | 75 | 50
FET Bias | [9 empty] | Fout | 0.018311 | 0.012207
Unit
Vcc | 9 | 9 | V
R1_series | 1000 | 1000 | Ω
VR | 1000 | 1000 | Ω |  |  |  | Connector
R_Shunt | 1000 | 1000 | Ω |  |  |  |  | V[Vac] | I[A] | Pin # | I_total[A]
V_Bias | 3 | 3 | V |  |  |  | HIF3H-xxPA-2.54DSA | 300 | 3 | 16 | 48
P_Shunt | 0.009 | 0.009 | W |  |  |  | HIF3H-xxDA-2.54DSA | 300 | 3 | 16 | 48
[7 empty] | MX5569 | 600 | 9 | 6 | 54
[7 empty] | MX35313 | 250 | 7 | 2 | 14
[7 empty] | BR1104C | 300 | 30 | 1 | 30
[7 empty] | BR950C | 250 | 18 | 1 | 18
[7 empty] | BR762C | 250 | 10 | 1 | 10
[7 empty] | MX5267 | 250 | 3 | 2 | 6

## FILTER
RF Input FILTER
RF Output FILTER

## Core
Core Specification
| CORE | Freq Range | Size[mm] |  |  | AL | Freq [MHz] | Turn calculation |  |  |  | L calculation
|  |  | OD | ID | Ht |  |  | Turn | L[uH] | Z[Ohm] | Length | L[uH] | Turn | Z[Ohm]
| FT114-61 | Wideband : 20~500MHz High-Q 0.2 - 10 MHz | 29 | 19 | 7.5 | 75 | 27.12 | 25 | 46.875 | 7983.45 | 25 | 20 | 16.3299 | 3406.272
| FT82-61 |  | 21 | 13.2 | 6.35 | 79 | 27.12 | 25 | 49.375 | 8409.234 |  | 20 | 15.9111 | 3406.272
| FT140-43 | Wideband : 5~400MHz PWR Trans : 0.5 - 30 MHz | 35.55 | 23 | 12.7 | 885 | 13.56 | 5 | 22.125 | 1884.0942 |  | 20 | 4.7538 | 1703.136
| FT114-43 |  | 29 | 19 | 7.5 | 510 | 27.12 | 25 | 318.75 | 54287.46 |  | 20 | 6.2622 | 3406.272
| FT114-43 |  | 29 | 19 | 7.5 | 510 | 13.56 | 25 | 318.75 | 27143.73 |  | 20 | 6.2622 | 1703.136
| FT114-43 |  | 29 | 19 | 7.5 | 510 | 1 | 25 | 318.75 | 2001.75 |  | 20 | 6.2622 | 125.6 | Test용 Coupler
| FT82-43 |  | 21 | 13.2 | 6.35 | 470 | 27.12 | 25 | 293.75 | 50029.62 |  | 20 | 6.5233 | 3406.272
| T106-6 | High-Q : 3 ~ 40MHz Cxx-Fx | 26.9 | 14.5 | 11.1 | 11.6 | 27.12 | 25 | 7.25 | 1234.7736 |  | 1 | 9.2848 | 170.3136 | Test용 Coupler
| T106-6 |  | 26.9 | 14.5 | 11.1 | 11.6 | 13.56 | 9 | 0.9396 | 80.0133 |  | 1 | 9.2848 | 85.1568
| T94-6 |  | 23.9 | 14.2 | 7.92 | 7 | 27.12 | 25 | 4.375 | 745.122 |  | 1 | 11.9523 | 170.3136
| T80-6 |  | 20.2 | 12.6 | 6.35 | 4.5 | 27.12 | 25 | 2.8125 | 479.007 |  | 1 | 14.9071 | 170.3136
| T50-6 |  | 12.7 | 7.7 | 4.83 | 4 | 13.56 | 7 | 0.196 | 16.6907 |  | 0.2 | 7.0711 | 17.0314
| T106-2 | High-Q : 250KHz ~ 10MHz Cxx-Ex | 26.9 | 14.5 | 11.1 | 13.5 | 27.12 | 25 | 8.4375 | 1437.021 |  | 1 | 8.6066 | 170.3136
| T94-2 |  | 23.9 | 14.2 | 7.92 | 8.4 | 27.12 | 25 | 5.25 | 894.1464 |  | 1 | 10.9109 | 170.3136
| T94-2 |  | 23.9 | 14.2 | 7.92 | 8.4 | 13.56 | 25 | 5.25 | 447.0732 |  | 1 | 10.9109 | 85.1568
| T80-2 |  | 20.2 | 12.6 | 6.35 | 5.5 | 27.12 | 25 | 3.4375 | 585.453 |  | 1 | 13.484 | 170.3136
| T50-2 |  | 12.7 | 7.7 | 4.83 | 4.9 | 27.12 | 25 | 3.0625 | 521.5854 |  | 0.1 | 4.5175 | 17.0314
| C23-E9 | 250kHz~10MHz | 23.06 | 13.97 | 9.53 | 9.2 | 27.12 | 25 | 5.75 | 979.3032 |  | 0.1 | 3.2969 | 17.0314
| C18-E6 |  | 17.73 | 9.4 | 6.35 | 7 | 27.12 | 25 | 4.375 | 745.122 |  | 0.1 | 3.7796 | 17.0314 | T68-2A
| C27-A11 | ~100kHz | 26.92 | 14.48 | 11.1 | 93 | 27.12 | 25 | 58.125 | 9899.478 |  | 0.1 | 1.037 | 17.0314 | T106-26
| C18-A6 |  | 17.53 | 9.4 | 6.35 | 58 | 27.12 | 25 | 36.25 | 6173.868 |  | 0.1 | 1.3131 | 17.0314 | T68-26A
| C21-A6 |  | 20.2 | 12.57 | 6.35 | 46 | 27.12 | 25 | 28.75 | 4896.516 |  | 0.1 | 1.4744 | 17.0314 | T80-26
|  |  | PCB | Ref | Core | Freq | Turn | L[uH] | Wire[Φ] | F[MHz] |  | Z[Ω]
|  |  | RF | L6 | FT140-43 | 5~500MHz | 5 | 24 | 1.6 | 27.12 |  | 4087.5264
|  |  | RF | L4,L5 | T50-6 | 3MHz ~ 40MHz | 6 | 0.5 | 1 | 27.12 |  | 85.1568
|  |  | RF | L7,L8,L9 | T106-6 | 3MHz ~ 40MHz | 9 | 1 | 1
|  |  | Coupler |  | FT82-43 | 5~500MHz | 25 | 290 | 0.45 | 27.12 |  | 49390.944
|  |  | Coupler | [5 empty] | 0.45 | 27.12 |  | 0
|  |  | RF | TX1 | RF-800 |  | 1 | 4 |  | 13.56 |  | 340.6272
|  |  | RF | TX2 - 1차 | RF-1000 |  | 2 | 25.6 |  | 13.56 |  | 2180.0141
|  |  | RF | TX2 - 2차 | RF-1000 |  | 1/2 | 3.3 |  | 13.56 |  | 281.0174
|  |  | RF | L6,L7,L8 | AIR - 20Φ |  | 5 | 0.5 | 1.6 | 13.56 |  | 42.5784
[7 empty] | 6 | 0.6 |  | 13.56 |  | 51.0941
[7 empty] | 8 | 1
[7 empty] | 10 | 1.4
Impedance Calculation
|  |  |  | Value | Freq | Z[Ω] |  |  | Value | Freq |  | Z[Ω]
|  |  | C[pF] | 1000 | 27.12 | 5.8715 |  | L[uH] | 20 | 27.12 |  | 3406.272
|  |  | C[nF] | 100 | 27.12 | 0.058715 |  | L[uH] | 100 | 27.12 |  | 17031.36
|  |  | C[nF] | 200 | 27.12 | 0.029358
|  |  | C[nF] | 1000 | 27.12 | 0.0058715

## COREf
| OD | ID | Ht | AL | L[uH] | Z[Ohm] | Turn
FT140-43 | 35.55 | 23 | 12.7 | 885 | 20 | 1704 | 4.8 | 5~500MHz
FT82-43 | 21 | 13.2 | 6.35 | 470 | 20 | 1704 | 6.5 | 5~500MHz
T106-6 | 26.9 | 14.5 | 11.1 | 11.6 | 1 | 85.2 | 9.3 | 3MHz ~ 40MHz
T106-2 | 26.9 | 14.5 | 11.1 | 13.5 | 1 | 85.2 | 8.6 | 250KHz ~ 10MHz
T50-6 | 12.7 | 7.7 | 4.83 | 4 | 0.1 | 8.5 | 5 | 3MHz ~ 40MHz
T50-6 | 12.7 | 7.7 | 4.83 | 4 | 0.25 | 21.3 | 8 | 3MHz ~ 40MHz
측정기가 기본적으로 0.3uH를 표시함, offset 감안하면 0.25uH로 유추 가능함
[10 empty] | 34.16 | 9.91
[10 empty] | 17.08 | 4.955

## Coupler
Long-wave SWR meter
Circuit
1) Transformer
100Ω 2W 두개를 사용하여 50Ω 4W 설계
turn ratio 1:25의 동일한 Transformer 2ea 사용
|  | RF-INPUT |  |  |  |  | coupling 1:25 |  | Load - 50Ω |  |  | Load - 25Ω
|  | P[W] | Vrms[V] | Vpp[V] | Irms[A] | Ipp[A] | Vrms[V] | Vpp[V] | P[W]-50 | Irms[A] | Ipp[A] | P[W]-25 | Irms[A] | Ipp[A]
|  | 10 | 22.3607 | 63.2551 | 0.44721 | 1.2651 | 0.89443 | 2.5302 | 0.016 | 0.017889 | 0.050604 | 0.032 | 0.035777 | 0.10121
|  | 100 | 70.7107 | 200.0302 | 1.4142 | 4.0006 | 2.8284 | 8.0012 | 0.16 | 0.056569 | 0.16002 | 0.32 | 0.11314 | 0.32005
|  | 300 | 122.4745 | 346.4625 | 2.4495 | 6.9292 | 4.899 | 13.8585 | 0.48 | 0.09798 | 0.27717 | 0.96 | 0.19596 | 0.55434
|  | 600 | 173.2051 | 489.9719 | 3.4641 | 9.7994 | 6.9282 | 19.5989 | 0.96 | 0.13856 | 0.39198 | 1.92 | 0.27713 | 0.78396
|  | 1000 | 223.6068 | 632.5511 | 4.4721 | 12.651 | 8.9443 | 25.302 | 1.6 | 0.17889 | 0.50604 | 3.2 | 0.35777 | 1.0121
R5025 0.67W - 4ea 사용
2차측 impedance가 Load impedance의 10배의 impedance를 가져야 한다.
|  | Core |  | FT140-43 | FT114-43 | FT82-43 | FT50-43 | FT114-43 | FT114-43 | UNIT
|  | Core OD |  | 35.55 | 29 | 21 | 12.7 | 29 | 29 | mm
|  | AL |  | 885 | 510 | 470 | 440 | 510 | 510
|  | uH for 25T |  | 553.125 | 318.75 | 293.75 | 275 | 318.75 | 318.75 | uH
|  | Start frequency |  | 0.3 | 0.3 | 0.3 | 0.3 | 27.12 | 13.65 | MHz
|  | 2차측 impedance |  | 1042.0875 | 600.525 | 553.425 | 518.1 | 54287.46 | 27323.8875 | Ω
|  | 2차측 권선 길이 |  | 1.226 | 0.78 | 0.78 | 0.78 | 0.78 | 0.78 | m
2차측 권선의 길이가 가장 짧은 파장의 1/10 이하 이어야 한다.
|  | End frequency |  | 30 | MHz
|  | 파장 |  | 10 | m
|  | 2차측 권선 최대 길이 |  | 1 | m
2) rectified circuit
VT - 1차측의 전류는 Load Z와 Trans Z로 계산해야 함.(Transformer 공식 사용하면 안됨)
|  | RF-INPUT |  |  |  |  | VT Turn ratio |  |  | 1차 | 25 | R1 |  | R2
[10 empty] | 2차 | 1 | R1[KΩ] | 1 | R2[KΩ] | 1
|  | P[W] | Vrms[V] | Vpp[V] | Irms[A] | Ipp[A] | Vrms[V] | Vpp[V] | P[W] | Irms[A] | Ipp[A] | Vrms[V] | P[W] | Vrms[V] | P[W] | Vp[V]
|  | 50 | 50 | 141.4427 | 1 | 2.8289 | 2 | 5.6577 | 0.08 | 0.04 | 0.11315 | 1 | 0.001 | 1 | 0.001 | 1.4144
|  | 100 | 70.7107 | 200.0302 | 1.4142 | 4.0006 | 2.8284 | 8.0012 | 0.16 | 0.056569 | 0.16002 | 1.4142 | 0.002 | 1.4142 | 0.002 | 2.0003
|  | 150 | 86.6025 | 244.986 | 1.7321 | 4.8997 | 3.4641 | 9.7994 | 0.24 | 0.069282 | 0.19599 | 1.7321 | 0.003 | 1.7321 | 0.003 | 2.4499
|  | 200 | 100 | 282.8854 | 2 | 5.6577 | 4 | 11.3154 | 0.32 | 0.08 | 0.22631 | 2 | 0.004 | 2 | 0.004 | 2.8289
|  | 250 | 111.8034 | 316.2755 | 2.2361 | 6.3255 | 4.4721 | 12.651 | 0.4 | 0.089443 | 0.25302 | 2.2361 | 0.005 | 2.2361 | 0.005 | 3.1628
|  | 300 | 122.4745 | 346.4625 | 2.4495 | 6.9292 | 4.899 | 13.8585 | 0.48 | 0.09798 | 0.27717 | 2.4495 | 0.006 | 2.4495 | 0.006 | 3.4646
|  | 350 | 132.2876 | 374.2223 | 2.6458 | 7.4844 | 5.2915 | 14.9689 | 0.56 | 0.10583 | 0.29938 | 2.6458 | 0.007 | 2.6458 | 0.007 | 3.7422
CT - 1차측의 전압은 Load Z와 Trans Z로 계산해야 함.(Transformer 공식 사용하면 안됨)
|  | RF-INPUT |  |  |  |  | CT Turn ratio |  |  | 1차 | 1 | R1 |  | R2
[10 empty] | 2차 | 25 | R1[KΩ] | 1 | R2[KΩ] | 1
|  | P[W] | Vrms[V] | Vpp[V] | Irms[A] | Ipp[A] | Irms[A] | Ipp[A] | Vrms[V] | Vpp[V] | P[W] | Vrms[V] | P[W] | Vrms[V] | P[W] | Vp[V]
|  | 50 | 50 | 141.4427 | 1 | 2.8289 | 0.04 | 0.11315 | 2 | 5.6577 | 0.08 | 1 | 0.001 | 1 | 0.001 | 1.4144
|  | 100 | 70.7107 | 200.0302 | 1.4142 | 4.0006 | 0.056569 | 0.16002 | 2.8284 | 8.0012 | 0.16 | 1.4142 | 0.002 | 1.4142 | 0.002 | 2.0003
|  | 150 | 86.6025 | 244.986 | 1.7321 | 4.8997 | 0.069282 | 0.19599 | 3.4641 | 9.7994 | 0.24 | 1.7321 | 0.003 | 1.7321 | 0.003 | 2.4499
|  | 200 | 100 | 282.8854 | 2 | 5.6577 | 0.08 | 0.22631 | 4 | 11.3154 | 0.32 | 2 | 0.004 | 2 | 0.004 | 2.8289
|  | 250 | 111.8034 | 316.2755 | 2.2361 | 6.3255 | 0.089443 | 0.25302 | 4.4721 | 12.651 | 0.4 | 2.2361 | 0.005 | 2.2361 | 0.005 | 3.1628
|  | 300 | 122.4745 | 346.4625 | 2.4495 | 6.9292 | 0.09798 | 0.27717 | 4.899 | 13.8585 | 0.48 | 2.4495 | 0.006 | 2.4495 | 0.006 | 3.4646
|  | 350 | 132.2876 | 374.2223 | 2.6458 | 7.4844 | 0.10583 | 0.29938 | 5.2915 | 14.9689 | 0.56 | 2.6458 | 0.007 | 2.6458 | 0.007 | 3.7422
RC Filter 설계
|  | R | 1 | 1 | 1 | KΩ
|  | C | 100 | 47 | 100000 | pF
|  | Fc | 1592.3567 | 3387.993 | 1.5924 | KHz
|  | τ | 0.1 | 0.047 | 100 | usec
Bypass Cap
|  | C | 20 | 68 | 100 | 100000 | pF
|  | Fc | 27120 | 13560 | 1000 | 0.8 | KHz
|  | Z | 293.5761 | 172.6918 | 1592.3567 | 1990.4459 | Ω

## Sheet4
2) rectified circuit
VT - 1차측의 전류는 Load Z와 Trans Z로 계산해야 함.(Transformer 공식 사용하면 안됨)
| RF-INPUT |  |  |  |  | VT Turn ratio |  |  | 1차 | 25 | R1 |  | R2
[9 empty] | 2차 | 1 | R1[KΩ] | 1 | R2[KΩ] | 1
| P[W] | Vrms[V] | Vpp[V] | Irms[A] | Ipp[A] | Vrms[V] | Vpp[V] | P[W] | Irms[A] | Ipp[A] | Vrms[V] | P[W] | Vrms[V] | P[W] | Vpp[V]
| 50 | 3.42 | 9.6747 | 14.6199 | 41.3575 | 0.1368 | 0.38699 | 0.00037428 | 0.002736 | 0.0077397 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| 100 | 70.7107 | 200.0302 | 1.4142 | 4.0006 | 2.8284 | 8.0012 | 0.16 | 0.056569 | 0.16002 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| 150 | 86.6025 | 244.986 | 1.7321 | 4.8997 | 3.4641 | 9.7994 | 0.24 | 0.069282 | 0.19599 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| 200 | 100 | 282.8854 | 2 | 5.6577 | 4 | 11.3154 | 0.32 | 0.08 | 0.22631 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| 250 | 111.8034 | 316.2755 | 2.2361 | 6.3255 | 4.4721 | 12.651 | 0.4 | 0.089443 | 0.25302 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| 300 | 122.4745 | 346.4625 | 2.4495 | 6.9292 | 4.899 | 13.8585 | 0.48 | 0.09798 | 0.27717 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| 350 | 132.2876 | 374.2223 | 2.6458 | 7.4844 | 5.2915 | 14.9689 | 0.56 | 0.10583 | 0.29938 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
CT - 1차측의 전압은 Load Z와 Trans Z로 계산해야 함.(Transformer 공식 사용하면 안됨)
| RF-INPUT |  |  |  |  | CT Turn ratio |  |  | 1차 | 1 | R1 |  | R2
[9 empty] | 2차 | 25 | R1[KΩ] | 1 | R2[KΩ] | 1
| Load | Vrms[V] | Vpp[V] | Irms[A] | Ipp[A] | Irms[A] | Ipp[A] | Vrms[V] | Vpp[V] | P[W] | Vrms[V] | P[W] | Vrms[V] | P[W] | Vpp[V]
| 50 | 3.42 | 9.6747 | 0.0684 | 0.19349 | 0.002736 | 0.0077397 | 0.1368 | 0.38699 | 0.00037428 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| 100 | 70.7107 | 200.0302 | 1.4142 | 4.0006 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| 150 | 86.6025 | 244.986 | 1.7321 | 4.8997 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| 200 | 100 | 282.8854 | 2 | 5.6577 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| 250 | 111.8034 | 316.2755 | 2.2361 | 6.3255 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| 300 | 122.4745 | 346.4625 | 2.4495 | 6.9292 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| 350 | 132.2876 | 374.2223 | 2.6458 | 7.4844 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!

## HeatSink
YS 120V80
YS 84V84A

## 배치
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
RF PCB
PWR PCB
Matcher
CPU PCB